## Feuil1
| ventilateurs MHRA | Your project
ventilation | must have Continuous Mandatory Ventilation Mandatory mode must be either (ideally) Pressure Regulated Volume Control, or
(ideally) Pressure Regulated Volume Control, or
pressure controlled ventilation (PCV) or
minimally a volume controlled ventilation (VCV).
PRVC/Pressure Controlled - a set pressure is delivered for the period of inspiration and the volume achieved is measured and displayed. Ideally PRVC, an adaptive mode where the tidal volume is set and the lowest possible pressure is delivered to achieve this volume. Pressure Control Ventilation where the user has to provide the adaptive control to achieve tidal volume is only acceptable if the tidal volume delivered is clearly displayed and the user can set patient specific upper and lower tidal volume alarms to alert to the need to adjust the pressure.
Volume Control – the user sets a tidal volume and respiratory rate. The tidal volume is delivered during the inspiratory period. Acceptable only if additional pressure limiting controls are available, see Inspiratory Pressure section.
Should have a spontaneous breathing pressure support mode for those patients breathing to some extent themselves, e.g. BIPAP or SIMV-PC. The user sets an inspiratory pressure and an expiratory pressure. The ventilator can sense when a patient starts to breathe in and apply the inspiratory pressure, then sense when the patient starts to breathe out and apply the expiratory pressure (this pressure is still positive but lower than the inspiratory pressure).
If a pressure support mode is provided the RMVS must failsafe automatically onto mandatory ventilation if the patient stops breathing in this mode.
Inspiratory airway pressure | Plateau pressure should be adjusted to achieve volume and be limited to 35 cmH2O by default. It is acceptable if an option to increase this to 70 cmH2O in exceptional circumstances is provided. This must require a positive decision and action by the user
Peak pressure should be no more than 2 cmH2O greater than plateau pressure.
If volume control ventilation is used, the user must be able to set inspiratory airway pressure limit in the range at least 15 – 40 cmH2O in at least increments of 5 cmH2O.
There must be a mechanical failsafe valve that opens at 80 cmH2O.
Positive End Expiratory Pressure (PEEP) | Range 5-20 cm H2O adjustable in 5 cmH2O increments.
Patient breathing system must remain pressurised to at least the PEEP level setting at all times.
Inspiratory:Expiratory ratio | 1:2.0 (i.e. expiration lasts twice as long as inspiration).
Optionally adjustable in the range 1:1 – 1:3.
Respiratory Rate | Range 10 – 30 breaths per minute in increments of 2 (only in mandatory mode) can be set by the user.
Tidal Volume (Vt) setting, | Must have at least one setting of 400ml +/- 10 ml.
Ideally 350ml and 450 ml options.
Optionally Range 250 – 600 ml in steps of 50ml.
Even more optionally up to 800 ml.
Incoming Gas Supply. | All gas connectors and hoses must comply with BS EN ISO 5359:2014+A1:2017, ISO 5359:2014/AMD 1:2017 and BS 2050: 1978 Electrical Conductivity.
Must connect to wall pipeline oxygen supply via compatibe probes. If hose not permanently fixed to machine, then must connect with NIST (Non-Interchangeable Screw Thread to ISO 18082:2014/AMD 1:2017). Oxygen pipeline pressure is approximately 3.7 – 4.5 bar.
Oxygen supply from wall outlets outside of ICU and theatres is limited to approximately 6-10 lpm (HTM_02-01_Part_A) and so provision for a gas reservoir will be required to manage peak inspiratory flow rates of up to 100 lpm
Average oxygen consumption must be no more than 3 lpm. This may be allowed to increase as greater certainty is gained over oxygen supply.
An RMVS may include an oxygen concentrator as the source of oxygen. Note these will typically be limited to 10 lpm 96% oxygen.
Electricity Supply. | Should connect to hospital electrical circuit.
Must be PAT tested to the adapted IEC 60601, IEC 62353 standards
If electricity is required for functioning, RMVS must have a battery backup of at least 20 minutes in case of mains electricity failure.
Could utilise hot swappable batteries so that it can be run on battery supply for an extended period, e.g. 2 hours for within hospital transfer.
Must avoid harmful RF or EM emissions that could interfere with other critical machinery.
Gas supply to patient. | User must be able to control inspired oxygen proportion (FiO2). The percentage of oxygen in the gas being breathed in by the patient. Room air is 21% oxygen.
At least (50% or 60%) and 100% options
Ideally variable between 30 – 100 % in 10% steps.
Patient breathing system connections: the ventilator must present 22mm outside diameter (OD) ‘male’ standard connectors to ISO 5356-1:2015 on both outlet and inlet ports for connection to user supplied 22mm ‘female’ connectors on the breathing system. These must be rigid and robust (not plastic) and separated by a minimum of 10 cm between centres to accommodate filter HMEs.
All elements in the gas pathway must meet biological safety and lowpressure oxygen safety standards, especially to minimise risk of fire or contamination of the patient’s airway.
Infection control | All parts coming into contact with the patient’s breath must be either disposable or able to be decontaminated between patients.
All working components of the device must be contained within an impermeable casing.
All external surfaces must be cleanable in the likely event that they get respiratory secretions or blood splatter on them. Cleaning would be by healthcare workers manually wiping using an approved surface wipe with disinfectant or cloths and approved surface cleaning liquid.
There will be a separately sourced HMEF-bacterial-viral filter between the machine and patient which may impact on resistance within the system, which may need to be accounted for with some designs. The pressure being delivered to the patient is the specified pressure. If the filter has a resistance of, say 2 cmH2O at 30 lpm, the ventilator needs to output 37 cmH2O to achieve a set 35 cmH2O at the patient.
Could include facility for hot water humidifier to be included in breathing system
Monitoring and Alarms | Must alarm at:
a. Gas or electricity supply failure.
b. Machine switched off while in mandatory ventilation mode.
c. Inspiratory airway pressure exceeded.
d. Inspiratory and PEEP pressure not achieved (equivalent to disconnection alarm).
e. Tidal volume not achieved or exceeded.
Monitoring
Must show the current settings of tidal volume, frequency, PEEP, FiO2, ventilation mode.
Must show the actual current airway pressure
Should show the achieved tidal volume, breathing rate, PEEP, andFiO2.
If pressure support mode is provided there must be real time confirmation of each patient breath and an alarm if below acceptable range.
Could provide CO2 monitoring.
Biological safety The authoritative standard covering this area is ISO 18562-1:2017 “Biocompatibility evaluation of breathing gas pathways in healthcare applications. Evaluation and testing within a risk management process”. | 1. Materials of Construction (raw materials)
A) The chosen material must be reasonably pure and simple in nature (minimise the use of additives where possible)
B) For components requiring flexibility avoid the use of materials requiring plasticisers. Good candidates are those materials that belong to the polyolefin family, examples include polyethylene and polypropylene
C) For structural components materials such as polycarbonate or Acrylonitrile butadiene styrene (ABS) should be used without additives, although reinforcement with glass fibre would be acceptable
D) Polyvinyl chloride (PVC) must be avoided in the patient gas pathway
E) PVC should be avoided elsewhere
2. Manufacturing process (risk from contaminants)
A) Mould release agents used within extrusion or injection moulding techniques may be required in setting up the machine, they should not be needed once a process is in full scale production
B) Approximately, the first 20 or so items in an injection moulding production run should be discarded to minimise risk from contamination with mould release agents
C) Extrusion and moulding techniques are comparatively simple and well controlled; therefore, ventilators will not be required to be manufactured within cleanroom specifications
D) Manufacture in a reasonably clean room and protection of components and products from contamination should suffice
E) If A-D is followed, chemical or particulate testing of the air coming out of the breathing circuit should not be necessary
3. Hazard Mitigation
A) Particulate matter: solid particles suspended in a gas- Particulate matter emissions are not of significant concern if the manufacturing process is adequately controlled as per the above criteria
B) Volatile organic compound (VOC): organic compound whose boiling point is in the range of 50⁰C to 260⁰C- Risk of exposure to VOCs can be minimised through the appropriate choice of materials as set out in section 1
C) Leachable substances (in condensate): chemical removed from the medical device by the action of water, other liquids or other gases related to the use of the medical device- insure a HME filter is used between the ventilator and breathing system.
Miscellaneous | Must be reliable. RMVS must be capable of continuous operation (100% duty cycle) for 14 days.
Should be capable of operation continuously for more than 14 days. The expected durability must be specified.
It must be intuitive to use for qualified medical personnel, but these may not be specialists in ventilator use.
a. Must not require more than 30 minutes training for a doctor with some experience of ventilator use.
b. Must include Instructions for Use.
c. Ideally instructions for use should be built into the labelling of the ventilator, e.g. with ‘connect this to wall’ etc.
d. Must include clear labelling of all critical functions and controls using standard terms, pictograms and colours that will be readily recognised by UK healthcare staff.
Must have transparent design, supply chain, manufacture, quality assurance and testing processes that are of sufficient quality to enable FAGG officials to deem appropriate for usage in exceptional circumstances.
Must be made from materials and parts readily available
Testing | compliance with the essential safety standards of iso 80601-2-12 and 60601-1 must be demonstrated for patient safety.
It is not anticipated that devices will be CE marked and approval by the FAGG will be through the “Exceptional use of non-CE marked medical devices” route.
When the current emergency has passed these devices will NOT be usable for routine care unless they have been CE marked through the Medical Device Regulations. The device must display a prominent indelible label to this effect.
Usability testing at both prototype and final production stages will be required. This should be done as a short Formative Usability Test to ISO62366 (this can be done in a day) in a realistic environment if possible.
The user must be able to instantly see the settings selected and be able to easily operate all controls while dressed in protective gear. They may be required to remain so clothed and operating the ventilator for a number of hours without breaks.
Software safety | Software in a high-risk device like RMVS will almost certainly have the capability to cause serious injury or death if risk control measures are not adequately implemented. See appendix C for details of minimal necessary risk control measures.