Draft MSOG document version 17092002			prepared by Philippe Bauwin

GUIDANCE NOTES FOR MANUFACTURERS OF CLASS I MEDICAL DEVICES

INTRODUCTION

These guidelines are intended for  medical devices manufacturers wishing to put Class I devices on the European market.

The 93/42/EEC European Directive regulates medical devices. Manufacturers should refer to their national legislation transposing  this directive.

The directive specifies essential requirements which must be met before any device can be placed on the market. It specifies also requirements for the assessment of clinical investigations protocols and the evaluation of any adverse incidents that occur.

Medical devices are classified in different risk classes as a function of their use and describes the conformity assessment procedure to the provisions of the directive for the various classes.

Class I medical devices are devices with a  relatively low risk.

Medical devices must carry a CE mark of conformity when they are placed on the market.

In addition to this guideline the manufacturer should refer to the national legislation and guidelines which are usually available on the national authorities websites .

The European Directive , the European guidelines and the list of related European harmonised standards can be found on the European Commission's web site

http://www.europa.eu.int/comm/enterprise/medical_devices/index.htm

# STEPS TO BE FOLLOWED

1) VERIFY IF YOUR PRODUCT IS A MEDICAL DEVICE AND IF IT IS CLASS I

'medical device' means any instrument, apparatus, appliance, material or other article, whether used alone or in combination, including the software necessary for its proper application intended by the manufacturer to be used for human beings for the purpose of: 
- diagnosis, prevention, monitoring, treatment or alleviation of disease, 
- diagnosis, monitoring, treatment, alleviation of or compensation for an injury or handicap, 
- investigation, replacement or modification of the anatomy or of a physiological process, 
- control of conception, 
and which does not achieve its principal intended action in or on the human body by pharmacological, immunological or metabolic means, but which may be assisted in its function by such means;

In Appendix A you will find examples of Class I medical devices.

You will find in Appendix B examples of products which are not medical devices but individual protection devices ,  cosmetic products or  other

Classification rules can be found in Annex IX of the Directive .

The Guideline : classification of medical devices Meddev 2.4/1 gives additional guidance and many examples

The Guidelines Meddev 2.1/3 and 2.1/4 explain the demarcation with other European Directives

If you have still doubts contact your National Competent Authority

2) IF YOUR PRODUCT IS A MEDICAL DEVICE AND OF CLASS I  FOLLOW THE REQUESTED CONFORMITY ASSESMENT PROCEDURE

Article 11 point 5 of the Directive says:

In the case of devices falling within Class I, other than devices which are custom-made or intended for clinical investigations, the manufacturer shall, in order to affix the CE marking, follow the procedure referred to in Annex VII and draw up the EC declaration of conformity required before placing the device on the market.

What this means will be explained below

Remark:

Don’t forget to take in account art 1 point 5 of the directive that lists the type of products to which the Directive does not apply (for example devices incorporating stable human blood derivatives that are submitted to an other directive) and the special classification rules under point 4 annex IX.

For example , rule 13 says: all devices incorporating , as an integral part, a substance which if used separately , can be considered to be a medicinal product , as defined in article 1 of Directive 65/65/EEC , and which is liable to act on the human body with action ancillary to that of the devices , are in class III.

Rule 17 classifies most devices manufactured utilizing animal tissues or derivatives in class III

3) BEFORE PUTTING THE PRODUCT ON THE MARKET NOTIFY YOUR NATIONAL COMPETENT AUTHORITY
. 
Article 14 of the Directive says:

Registration of persons responsible for placing devices on the market

- Any manufacturer who, under his own name, places Class I devices on the market in accordance with the procedures referred to in Article 11 (5) ……………………shall inform the competent authorities of the Member State in which he has his registered place of business of the address of the registered place of business and the description of the devices concerned.

If your place of business is outside the Member States follow this procedure

2. Where a manufacturer who places devices referred to in paragraph 1 on the market under his own name does not have a registered place of business in a Member State, he shall designate the person(s) responsible for marketing them who is (are) established in the Community. These persons shall inform the competent authorities of the Member State in which they have their registered place of business of the address of the registered place of business and the category of devices concerned.

Refer to your national authority for the notification process in your country.

REMARK:

It is important to realize that this notification is not a simple formality. Before notifying the manufacturer shall perform all activities needed to ensure that the product is safe and performs as requested and prepare a technical documentation that must allow assessment of the conformity of the product with the requirements of  the directive. This documentation shall be made available on request to the competent authorities. Then he will make a declaration of conformity that specifies that the product concerned meets the provisions of the Directive that applies to it. Only then has he the right to apply the CE mark on the product and notify the national authority. There are post-marketing obligations such as the notification of incidents . Authorities are responsible for market surveillance and in this context will monitor the activities of their Class I manufacturers.

## CONFORMITY ASSESMENT PROCEDURE

## ACCORDING ANNEX VII OF THE DIRECTIVE

## 1) THE TECHNICAL DOCUMENTATION

  - GENERAL INFORMATION:

The manufacturer must prepare the technical documentation described below. The manufacturer or his authorized representative established in the Community must make this documentation , including the declaration of conformity, available to the national authorities for inspection purposes for a period ending at least five years after the last product has been manufactured.

The technical documentation must allow assessment of the conformity of the products with the requirements of the directive.

The devices must meet the essential requirements set out in Annex I of the Directive which apply to them, taking in account of the intended purpose of the devices concerned.

The documentation must include in particular:

- a general description of the product, including any variants planned, 
- design drawings, methods of manufacture envisaged and diagrams of components, sub-assemblies, circuits, etc., 
- the descriptions and explanations necessary to understand the abovementioned drawings and diagrams and the operations of the product, 
- the results of the risk analysis and a list of the standards referred to in Article 5 of the Directive, applied in full or in part, and descriptions of the solutions adopted to meet the essential requirements of the Directive if the standards referred to in Article 5 have not been applied in full, 
- in the case of products placed on the market in a sterile condition, description of the methods used, 
- the results of the design calculations and of the inspections carried out, etc.; if the device is to be connected to other device(s) in order to operate as intended, proof must be provided that it conforms to the essential requirements when connected to any such device(s) having the characteristics specified by the manufacturer, 
- the test reports and, where appropriate, clinical data in accordance with Annex X, 
- the label and instructions for use.

Member States shall presume compliance with the essential requirements  in respect of devices which are in conformity with the relevant national standards adopted pursuant to the harmonized standards the references of which have been published in the Official Journal of the European Communities; Member States shall publish the references of such national standards.

Reference to harmonized standards also includes the monographs of the European Pharmacopoeia notably on surgical sutures and on interaction between medicinal products and materials used in devices containing such medicinal products, the references of which have been published in the Official Journal of the European Communities.

  - DETAILED INFORMATION  ON THE NORMAL CONTENT OF THE TECHNICAL DOCUMENTATION

Examples are given of what the documentation would normally  include. These are not exhaustive and each case has to be examined individually. Data available from the literature , from component suppliers or preliminary documented clinical experience can in some cases be used. Depending on the type of product some tests will not be required. This can be determined by a proper risk analysis.

The description of the product:

A general description of the product including any variants (for example names , model numbers, size). Photographies or drawings will be useful.

Design documentation such as drawings, calculation results and design verification . Tests would include for example process validation data, performance tests , validation of the shelf life  as appropriate…

Component documentation: raw material composition and biocompatibility data , components diagrams , incoming quality control procedures…

Intermediate products and sub-assemblies documentation: specifications, including appropriate drawings and/or master patterns, circuits and formulation specifications, manufacturing methods , process validation data and any quality control procedures…

Final product documentation : specifications, including appropriate drawings, and :or master patterns, circuits and formulation specifications, relevant manufacturing methods  including sterilisation if sterile product, process validation data and any quality control procedures…

Appropriate descriptions and explanations must be provided to understand the above mentioned  data, procedures , drawings and diagrams and the operations of the product

Packaging and labelling documentation : packaging specifications and copies of all labels and instructions for use..

Detailed information regarding the information to be supplied with the product is given in the Annex I of the Directive and in the standard on labelling

Risk analysis: the result of risk analysis to review  any risk associated with the use of the products.

It must be done in a systematic way using the appropriate standard (EN 1441 Medical devices-risk analysis) .The solutions adopted to eliminate as far as possible the risks must be described.

At the end the balance risk/benefit must be evaluated.

Typical risks are for example:

toxicity of the materials used (for example the risk of skin irritation by dressing material)

electric shock or electromagnetic interference in case of powered devices

List of standards used: products manufactured according the harmonised standards benefit of a presumption of conformity to the essential requirements. The manufacturer will indicate here the standards he used and shall describe the solutions adopted to meet the essential requirements if these standards have not been applied in full (for example if he uses his own standards or alternative standards)

There are horizontal standards like

risk analysis (EN1441)

sterilization (EN550 serial)

biocompatibility (EN 10993-)

electrical safety (EN 60601-)

quality systems for medical devices manufacturers (EN 46000 serial)

and product group standards

For example the EN 10993-1 Biological evaluation of medical devices-Part 1 Evaluation and testing  includes a flow chart to aid in a systematic approach to the biological evaluation of your device that will help you to determine if tests are needed or not (tests may not be necessary if sufficient data are available) and to select the appropriate tests . Typical tests for class I devices are cytotoxicity , sensitization and irritation or intracutaneous reactivity.

Clinical data: many class I devices will not require a special clinical investigation to establish data on performance and safety or side effects. For products which have been established for a number of years and those which are modifications of such products , it is likely that a compilation and review of existing clinical experience would be sufficient to cover this requirement.

However all manufacturers should review the intended use of the product and any medical claims that are being made to ensure that they have both adequate supporting test results and records of relevant experience

Note that if a clinical investigation is required to justify use of a device , then the Competent Authority requires advance notification of the proposal

Records: manufacturing and test records to show compliance with the defined procedures and specifications

## Special requirements

In case of products placed on the market in sterile condition and devices with a measuring function , the manufacturer must also observe one of the procedures referred to in Annex IV , V , or VI of the Directive. This requires the intervention of a notified body.

### Class I devices are considered to have a measuring function if they measure a parameter such as a volume or a physiological parameter and display or indicate its value in a unit of measurement (example: urine bags or thermometers)

The intervention by the notified body is limited to: 
- in the case of products placed on the market in sterile condition, only the aspects of manufacture concerned with securing and maintaining sterile conditions, 
- in the case of devices with a measuring function, only the aspects of manufacture concerned with the conformity of the products with the metrological requirements.

2) THE EC DECLARATION OF CONFORMITY

The EC declaration of conformity is the procedure whereby the manufacturer or his authorized representative established in the Community who fulfils the obligations imposed by Section 2 of the Annex VII of the directive (the technical documentation) and, in the case of products placed on the market in a sterile condition and devices with a measuring function, the obligations imposed by Section 5 of the Annex  (see above: special requirements ) ensures and declares that the products concerned meet the provisions of this Directive which apply to them.

Where the Annex VII is applied in conjunction with the procedure referred to in Annex IV, V or VI of the directive (products placed on the market in a sterile condition and devices with a measuring function), the declaration of conformity referred to in the abovementioned Annexes forms a single declaration. As regards the declaration based on  Annex VII , the manufacturer must ensure and declare that the product design meets the provisions of this Directive which apply to it.

.
3)  CE MARKING

3.1. Devices, other than devices which are custom-made or intended for clinical investigations, considered to meet the essential requirements referred to in Article 3 of the directive must bear the CE marking of conformity when they are placed on the market.

3.2. The CE marking of conformity, as shown in Annex XII (see below), must appear in a visible, legible and indelible form on the device or its sterile pack, where practicable and appropriate, and on the instructions for use. Where applicable, the CE marking must also appear on the sales packaging. 
It shall be accompanied by the identification number of the notified body responsible for implementation of the procedures set out in Annexes IV, V and VI. (products placed on the market in sterile condition and devices with a measuring function)

As a consequence , Class I devices other than products placed on the market in sterile condition and devices with a measuring function will bear a CE mark without a notified body number

3.3. It is prohibited to affix marks or inscriptions which are likely to mislead third parties with regard to the meaning or the graphics of the CE marking. Any other mark may be affixed to the device, to the packaging or to the instruction leaflet accompanying the device provided that the visibility and legibility of the CE marking is not thereby reduced.

ANNEX XII 
CE MARKING OF CONFORMITY The CE conformity marking shall consist of the initials 'CE' taking the following form: 
- If the marking is reduced or enlarged the proportions given in the above graduated drawing must be respected. 
- The various components of the CE marking must have substantially the same vertical dimension, which may not be less than 5 mm. 
This minimum dimension may be waived for small-scale devices.

## OBLIGATIONS OF THE MANUFACTURER

The manufacturer or his authorized representative established in the Community must make the technical  documentation, including the declaration of conformity, available to the national authorities for inspection purposes for a period ending at least five years after the last product has been manufactured.

The manufacturer shall institute and keep up to date a systematic procedure to review experience gained from devices in the post-production phase and to implement appropriate means to apply any necessary corrective actions, taking account of the nature and risks in relation to the product.

He shall notify the competent authorities of the following incidents immediately on learning of them: 
(i) any malfunction or deterioration in the characteristics and/or performance of a device, as well as any inadequacy in the labelling or the instructions for use which might lead to or might have led to the death of a patient or user or to a serious deterioration in his state of health; 
(ii) any technical or medical reason connected with the characteristics on the performance of a device for the reasons referred to in subparagraph (i) leading to systematic recall of devices of the same type by the manufacturer.

INFORMATION SOURCES

http://www.europa.eu.int/comm/enterprise/medical_devices/index.htm

http://europa.eu.int/comm/enterprise/newapproach/legislation/nb/notified-bodies.htm

APPENDIX A

# EXAMPLES OF CLASS I MEDICAL DEVICES

1)Devices that either do not touch the patient or contact intact skin only

- body liquid collection devices intended to be used in such a way that a return flow is unlikely (e.g. to collect body wastes such as urine collection bottles ostomy pouches, incontinence pads or collectors used with wound drainage devices.) They may be connected to the patient by means of catheters and tubing.

- Devices used to immobilize body parts and/or apply force or compression on them (e.g. non-sterile dressings used to aid the healing of a sprain, plaster of Paris, cervical collars, gravity traction devices, compression hosiery).

- Devices intended in general for external patient support (e.g. hospital beds, patient hoists, walking aids, wheelchairs, stretchers, dental patient chairs).

Corrective glasses, frames, stethoscopes for diagnosis, eye occlusion plasters, incision drapes, conductive gels, non-invasive electrodes, image intensifying screens, software intended for image processing.

2. Channelling or storing for eventual administration

- Devices that provide a simple channelling function, with gravity providing the force to transport the liquid, e.g. tubing used in gravity drips for saline solution and medicines.

- Devices intended to be used for a temporary containment or storage function such as cups and spoons specifically intended for administering medicines

- Syringes without needles

3. Devices in contact with injured skin

All non-invasive devices which come into contact with injured skin:

- are in Class I if they are intended to be used as a mechanical barrier, for compression or for absorption of exudates

Examples:

- Wound dressings, such as absorbent pads, {island dressings}, cotton wool, wound strips and gauze dressings to act as a barrier or to maintain the wound positionally or to absorb exudates from the wound

4. Devices invasive in body orifices

All invasive device with respect to body orifices, other than surgically invasive devices and which are not intended for connection to an active medical device:

are in Class I if they are intended for transient use,

Examples:

- handheld mirrors used in dentistry to aid in dental diagnosis and surgery, dental impression materials, tubes used for pumping the stomach, impression tray, enema devices, examination gloves and prostatic balloon dilation catheters

in case of short term use they are also class I if they are used in the oral cavity as far as the pharynx, in an ear canal up to the ear drum or in a nasal  cavity

Example:

- Dressings for nose bleeds, dentures removable by the patient

5. Surgically invasive devices for transient use

Reusable surgical instruments

Example:

- Scalpels, drill bits, saws, that are not intended for connection to an active device, and retractors forceps, excavators and chisels.

6. Some active devices

- Active diagnostic devices intended to illuminate the patient’s body in the visible spectrum such as examination lights or to optically view the body such as surgical microscopes.

- Devices intended in general for external patient support (e.g. hospital beds, patient hoists, walking aids, wheelchairs, stretchers, dental patient chairs).

- Active diagnostic devices intended for thermography.

- Active devices intended for recording, processing or viewing of diagnostic images.

- Dental curing lights.

APPENDIX B

Examples of products related to other directives and that are not medical devices

TOILETRY AND COSMETICS PRODUCTS, EXAMPLES:

- Tooth brushes, dental sticks, dental floss, unmedicated dental chewing gum

- Baby nappies, hygiene tampons, mattress protectors

- Contact lenses without corrective function intended to provide colour to the eyes

- Instruments for tattooing

- Deodorants for use with devices

- Wigs

PERSONAL PROTECTIVE EQUIPMENT, EXAMPLES:

- Self rescue apparatus

- Mouth guards

- Ionising radiation protective clothing

- Eye protecting visors

- Protective gloves (for example for use in a medical laboratory)

OTHER EQUIPMENT, EXAMPLES

- Acoustic signals at traffic lights

- Breathalysers, other blood or air alcohol measuring devices

- Non sterile home, occupational, recreational protective/safety apparel

- Non prescription sunglasses

- Consumer products aimed at comfort

- Products for sport or leisure

CHECK LIST CLASS I MANUFACTURERS

|  | YES | NO | REMARK |
| --- | --- | --- | --- |
| Preliminary |  |  |  |
| Status of medical device correct |  |  |  |
| Correctly classified in class I |  |  |  |
| No medicinal product as auxiliary substance |  |  |  |
| Technical documentation |  |  |  |
| Description of the product |  |  |  |
| Design documentation |  |  |  |
| Component documentation |  |  |  |
| Intermediate products/sub-assemblies documentation |  |  |  |
| Final product documentation |  |  |  |
| Packaging and labelling documentation |  |  |  |
| Risk analysis |  |  |  |
| List of standard used or alternative solutions |  |  |  |
| Clinical data |  |  |  |
| Manufacturing and test records |  |  |  |
| In case of sterile products or device with a measuring function , notified body certificate |  |  |  |
| Declaration of conformity |  |  |  |
| Correct declaration |  |  |  |
| EC marking |  |  |  |
| On the device |  |  |  |
| On the sterile pack |  |  |  |
| On the instructions for use |  |  |  |
| On sales packaging |  |  |  |
| Notified body number |  |  |  |
| Correct style |  |  |  |
| No additional marks |  |  |  |
| Post-marketing obligations |  |  |  |
| Technical documentation available |  |  |  |
| Procedures to review experience |  |  |  |
| Procedures to notify incidents |  |  |  |